Document	Bring up report
Project	Flexi Teer
PCB	Rev2.0
PCB SN	#1...#5

Authors	MD Anderson, / Dennis Yaskevich
Date	ongoing
Version	ongoing

## 1. Orientation
### 1.1 What is the purpose of this document
This document describes the initial bring up (verification) testing of the FlexiTeer rev.2 hardware. All of test performed and the issues detected during the testing must be reflected in this document.
### 1.2 Who should read this document
This is deeply technical document more inclined to electronics design. It mainly useful for the engineer who is designing the second revision of hardware to see at glance all of the issues the previous revision had.
### 1.3 Requisite documents
Table below contains the relevant design documentation and it’s current versions as of writing of this document.
Document name	Link to the document	Latest updated on
Schematic_Flexi-TEER_2023-02-18.pdf	GitHub	2023-02-18
PCB Layout	oshwlab.com	2023-02-18

### 1.4 Test environments
Currently 2 people are working on the project.
Test environment	Test engineer	Equipment available	Test to be performed
Boston Children Hospital	MD Anderson	To be completed	Shorts and resistance measurement
Hackerspace Krakow	Dennis Yaskevich	USB Microscope NB-MIKR-300	Visual inspection / Shorts and resistance measurement

## 2. Basic hardware testing
### 2.1 Visual inspection
Test equipment:
        - USB Microscope NB-MIKR-300

To be completed by Dennis

### 2.2 Shorts and resistance measurement
Test equipment:
  - Multimeter
  - lab power supply
Resistance measurement to be performed on to powered off unit according to the table below.
Signal name	Meas. point	Expected value	Actual value	Status
Q33_6	BT6 vs GND	>1k	Over Load
Q3_2	BT5 vs BT6	>1k	12.3M Ohms
VBAT	BT5 vs BT6	>1k	Same as above
VIN	R1 vs GND	>1k	Overload
+5V	TP31 vs GND	>1k	8.16k Ohms
+3V3	TP30 vs GND	>1k	8.55k Ohms
-5V	TP29 vs GND	>1k	230k Ohms