Document	Bring up report
Project	Flexi Teer
PCB	Rev2.0
PCB SN	#1...#5

Authors	MD Anderson, / Dennis Yaskevich
Date	ongoing
Version	ongoing

## 1. Orientation
### 1.1 What is the purpose of this document
This document describes the initial bring up (verification) testing of the FlexiTeer rev.2 hardware. All of test performed and the issues detected during the testing must be reflected in this document.
### 1.2 Who should read this document
This is deeply technical document more inclined to electronics design. It mainly useful for the engineer who is designing the second revision of hardware to see at glance all of the issues the previous revision had.
### 1.3 Requisite documents
Table below contains the relevant design documentation and it’s current versions as of writing of this document.
Document name	Link to the document	Latest updated on
Schematic_Flexi-TEER_2023-02-18.pdf	GitHub	2023-02-18
PCB Layout	oshwlab.com	2023-02-18

### 1.4 Test environments
Currently 2 people are working on the project.
Test environment	Test engineer	Equipment available	Test to be performed
Boston Children Hospital	MD Anderson	To be completed	Shorts and resistance measurement
Hackerspace Krakow	Dennis Yaskevich	USB Microscope NB-MIKR-300	Visual inspection / Shorts and resistance measurement

## 2. Basic hardware testing
### 2.1 Visual inspection
Test equipment:
        - USB Microscope NB-MIKR-300

To be completed by Dennis

### 2.2 Shorts and resistance measurement
Test equipment:
  - Multimeter
Resistance measurement to be performed on to powered off unit according to the table below.
Signal name	Meas. point	Expected value	Actual value	Status
Q33_6	BT6 vs GND	>1k	Over Load	OK
Q3_2	BT5 vs BT6	>1k	12.3M Ohms	OK
Q3_2	BT5 vs GND	>1k	15.2M Ohms	OK
VBAT	BT3 vs BT6	>1k	14M Ohms	OK
VBAT	BT3 vs GND	>1k	82k Ohms	OK
VIN	Left side of R1 vs GND	>1k	Overload	OK
+5V	TP31 vs GND	>1k	8.16k Ohms	OK
+3V3	TP30 vs GND	>1k	8.55k Ohms	OK
-5V	TP29 vs GND	>1k	230k Ohms	OK

## 2.3 Power up and voltage measurement

Test equipment:
  - Multimeter
  - lab power supply / 9V wall adapter

Issue E1 is detected when inserting the batteries. To be investigated by DY
Suggested temporary solution: Solder shorting 0 Ohm resistors R1 and R38. Wire can be soldered instead of this resistors. Apply 9V input power to the CON1 with 100mA current limit if possible.

If no overcurrent / smoke is detected after the power is applied, measure the following voltages:

Rail name	Meas. point	Expected value, V			Ref.	Actual value	Status
		Min.	Nom.	Max.
VIN	R1 vs GND	8	9	10		9.01	Ok
+5V	TP31 vs GND	4.5	5	5.5		4.98	Ok
+3V3	TP30 vs GND	3.0	3.3	3.5		3.26	Ok
-5V	TP29 vs GND	-5.5	-5	-4.5		-4.8	Ok
+1V65	TP27 vs GND		1.65			1.62	Ok
+2V5	TP25 vs GND		2.5			2.48	Ok

Set status as OK if measured voltage is within the expected values range.

Evaluating BiLevel converter and transform to make signals go between -5 and 5 volts to toggle MUX ICs.

Rail name	Meas. point	Expected value, V			Ref.	Actual value	Status
		Min.	Nom.	Max.
A in	A_3V3 vs GND		3.3			3.34	Ok
A shifted	A_5V vs GND		5			4.37	Ok
A -5_5	A_-5V_5V vs GND		-5 and 5			-4.75 and +4.35	Ok
MUX in	MUX_EN vs GND		3.3
MUX_EN -5_5	EN_-5V_5V vs GND		-5 and 5
OE enb	OE_LEVEL vs GND		3.3			2.77	High enough to work, but odd its so low.

Everything appears to be functional with this module.
## Summary
### Assembly issues
No	Issue description	Date/Author	Status
A1
A2

### Electronics issues

No	Issue description	Date/Author	Status
E1	After inserting the second battery the PCB trace starts burning	02.08.23 MD Anderson	To be investigated by DY
E2	Low humming coming from bilevel converter. Some outputs (C and A) stuck at 2.3V no matter what 3V3 corresponding input is.	15.08.23 MD Anderson	Swapped to different board and noise was not present and bilevel converter worked properly.